GOVERNMENT OF INDIA
MINISTRY OF CONSUMER AFFAIRS, FOOD & PUBLIC DISTRIBUTION
DEPARTMENT OF CONSUMER AFFAIRS — LEGAL METROLOGY DIVISION

TYPE EVALUATION TEST REPORT
Non-Automatic Weighing Instrument as per OIML R 76-1:2006 (E)

| Report Number: REP-2026-OIML-0C09BE | Date: 15-Sep-2026 |
| --- | --- |
| Session Number: TS-20260915-411F46 |  |
| Overall Result: NOT_TESTED | Standard: OIML R 76-1:2006 (E) |

## 1. Applicant & Manufacturer Information

| Applicant Name | WeighTech Metrology India Ltd. |
| --- | --- |
| Applicant Address | Plot 42, Electronics City, Bangalore, India |
| Manufacturer Name | WeighTech Metrology India Ltd. |
| Manufacturer Address | Plot 42, Electronics City, Bangalore, India |

## 2. Technical & Metrological Characteristics

| Model & Serial Number | MW-3000 / S/N: SN-E2E-TEST-EE44D4DB |
| --- | --- |
| Accuracy Class | Class III |
| Max & Min Capacity | Max = 3000.0 kg, Min = 20.0 kg |
| Scale Intervals (e, d, n) | e = 1.0 kg, d = 1.0 kg, n = 3000 |
| Load Receptor / Supports | Platform (4 support points) |
| Power & Temp Range | 230 V AC, 50 Hz ¦ -10.0°C to 40.0°C |

## 3. OIML R-76 Test Applicability & Coverage Matrix

| Code | Clause | Test Procedure | Applicability | Execution | Status |
| --- | --- | --- | --- | --- | --- |
| A.1 | Clause A.1 | Administrative Examination | APPLICABLE | NOT TESTED | PENDING |
| A.2 | Clause A.2 | Compare Construction with Documentation | APPLICABLE | NOT TESTED | PENDING |
| A.3 | Clause A.3 | Initial Examination | APPLICABLE | NOT TESTED | PENDING |
| A.4.2.1 | Clause 4.5.1 / A.4.2.1 | Range of Zero-Setting | APPLICABLE | NOT TESTED | PENDING |
| A.4.2.3 | Clause A.4.2.3 | Accuracy of Zero-Setting | APPLICABLE | NOT TESTED | PENDING |
| A.4.4.1 | Clause A.4.4.1 & A.4.4.3 | Weighing Performance Test | APPLICABLE | NOT TESTED | PENDING |
| A.4.7 | Clause A.4.7 | Eccentricity Test | APPLICABLE | NOT TESTED | PENDING |
| A.4.8 | Clause A.4.8 | Discrimination Test | APPLICABLE | NOT TESTED | PENDING |
| A.4.10 | Clause A.4.10 | Repeatability Test | APPLICABLE | NOT TESTED | PENDING |
| A.4.11.1 | Clause A.4.11.1 | Creep Test | APPLICABLE | NOT TESTED | PENDING |
| A.4.11.2 | Clause A.4.11.2 | Zero Return Test | APPLICABLE | NOT TESTED | PENDING |
| A.5.1 | Clause A.5.1 | Tilting Test | APPLICABLE | NOT TESTED | PENDING |
| A.5.3.1 | Clause A.5.3.1 | Static Temperatures Test | APPLICABLE | NOT TESTED | PENDING |
| A.5.3.2 | Clause A.5.3.2 | Temperature Effect on No-Load Indication | APPLICABLE | NOT TESTED | PENDING |
| A.5.4 | Clause A.5.4 | Voltage Variations Test | APPLICABLE | NOT TESTED | PENDING |
| A.6 | Clause A.6 | Endurance Test | APPLICABLE | NOT TESTED | PENDING |
| B.2 | Clause B.2 | Damp Heat, Steady State | APPLICABLE | NOT TESTED | PENDING |
| B.3 | Clause B.3 | Performance Tests for Disturbances (EMC / ESD / Dips) | APPLICABLE | NOT TESTED | PENDING |
| B.4 | Clause B.4 | Span Stability Test | APPLICABLE | NOT TESTED | PENDING |
| Annex G | Annex G (Cl. G.1 - G.4) | Software-Controlled Instrument Examination | APPLICABLE | NOT TESTED | PENDING |

## 4. Summary of Executed Test Results & Metrological Compliance

| Clause | Test Description | Observed Result | Permissible Limit | Result |
| --- | --- | --- | --- | --- |

5. Review & Approval Signatures

| Evaluated by: ___________________ Test Engineer Date: 15-Sep-2026 | Reviewed by: ___________________ Technical Reviewer Status: READY FOR REVIEW (Pending Review) | Approved by: ___________________ Director / Lab In-Charge Status: Pending Final Approval |
| --- | --- | --- |

DISCLAIMER: This prototype test report is generated by METRIREPORT for SIH 2026. It is not an official model approval certificate under the Legal Metrology Act.