GOVERNMENT OF INDIA
MINISTRY OF CONSUMER AFFAIRS, FOOD & PUBLIC DISTRIBUTION
DEPARTMENT OF CONSUMER AFFAIRS — LEGAL METROLOGY DIVISION

TYPE EVALUATION TEST REPORT
Non-Automatic Weighing Instrument as per OIML R 76-1:2006 (E)

| Report Number: REP-2026-OIML-001 | Date: 14-Sep-2026 |
| --- | --- |
| Session Number: TS-2026-001-MW3000 | Status: FINALIZED (v1.0) |
| Overall Result: PASS | Standard: OIML R 76-1:2006 (E) |

## 1. Applicant & Manufacturer Information

| Applicant Name | Demo WeighTech Pvt. Ltd. |
| --- | --- |
| Applicant Address | Plot 45, Okhla Industrial Area Phase III, New Delhi 110020, India |
| Manufacturer Name | Demo WeighTech Pvt. Ltd. |
| Manufacturer Address | Plot 45, Okhla Industrial Area Phase III, New Delhi 110020, India |

## 2. Technical & Metrological Characteristics

| Model & Serial Number | MW-3000 / S/N: MW3000-2026-IND-001 |
| --- | --- |
| Accuracy Class | Class III |
| Max & Min Capacity | Max = 3000.0 kg, Min = 20.0 kg |
| Scale Intervals (e, d, n) | e = 1.0 kg, d = 1.0 kg, n = 3000 |
| Load Receptor / Supports | Platform (4 support points) |
| Power & Temp Range | 230 V AC, 50 Hz ¦ -10.0°C to 40.0°C |

## 3. Summary of Test Results & Compliance

| Clause | Test Description | Observed Result | Permissible Limit | Result |
| --- | --- | --- | --- | --- |
| Clause A.4.2.3 | Accuracy of Zero-Setting | E0 = +0.050 kg | Limit = ±0.250 kg (0.25e) | PASS |
| Clause A.4.4.1 & A.4.4.3 | Weighing Performance Test | 5 points evaluated (Max ¦Ec¦ = 0.200 kg) | Tiered MPE as per Table 6 for Class III | PASS |
| Clause A.4.7 | Eccentricity Test | Max eccentricity error = 0.090 kg | Limit = ±1.000 kg | PASS |
| Clause A.4.10 & 3.6.1 | Repeatability Test | ΔE (Max-Min) = 0.030 kg | Limit = 1.000 kg | PASS |

4. Review & Approval Signatures

| Evaluated by: ___________________ Test Engineer Date: 14-Sep-2026 | Reviewed by: ___________________ Technical Reviewer Status: FINALIZED | Approved by: ___________________ Director / Lab In-Charge Status: FINALIZED / SEALED Date: 12-Sep-2026 |
| --- | --- | --- |

DISCLAIMER: This prototype test report is generated by METRIREPORT for SIH 2026. It is not an official model approval certificate under the Legal Metrology Act.